## Disclaimer
API publications necessarily address problems of a general nature. With respect to particular circumstances, local, state, and federal laws and regulations should be reviewed. Neither API nor any of API’s employees, subcontractors, consultants, committees, or other assignees make any warranty or representation, either express or implied, with respect to the accuracy, completeness, or usefulness of the information contained herein, or assume any liability or responsibility for any use, or the results of such use, of any information or process disclosed in this publication. Neither API nor any of API’s employees, subcontractors, consultants, or other assignees represent that use of this publication would not infringe upon privately owned rights. API publications may be used by anyone desiring to do so. Every effort has been made by the Institute to assure the accuracy and reliability of the data contained in them; however, the Institute makes no representation, warranty, or guarantee in connection with this publication and hereby expressly disclaims any liability or responsibility for loss or damage resulting from its use or for the violation of any authorities having jurisdiction with which this publication may conflict. API publications are published to facilitate the broad availability of proven, sound engineering and operating practices. These publications are not intended to obviate the need for applying sound engineering judgment regarding when and where these publications should be utilized. The formulation and publication of API publications is not intended in any way to inhibit anyone from using any other practices. Any manufacturer marking equipment or materials in conformance with the marking requirements of an API standard is solely responsible for complying with all the applicable requirements of that standard. API does not represent, warrant, or guarantee that such products do in fact conform to the applicable API standard.

## Revision
Date | Revision | Comment
4533200% | 0 | First Issue (GP)

## Installation Info&Instructions
API MPMS 10.10 Annex - WCA Installation General Information (to be completed for each Product Group)
Select or fill out each highlighted cell below
Location | Station XYZ
Installation Type | Insertion
Type of Sampling System | In Line
Line Size (inches) | 12"
Service | LACT
WCA Technology | Microwave
WCA ID | WCA 1
Sampling System ID | Metering System ABC Sampling
Product Group | WTI
General comments here (if any)
Instructions & General Notes
Annex B - Performance Acceptance Test
1. Select or fill out each highlighted cell
2. This data set is for FWA WCA comparison against the Water Injection Test
3. Please complete the form for pass as well as fail scenarios
4. Use the acceptance criteria in API Ch. 8.2 for pass or fail
Annex C - Ongoing Performance Test
1. Select or fill out each highlighted cell
2. This data set is only for FWA WCA comparison against the automatic composite sample (not manual samples)
3. Please complete the form for pass as well as fail scenarios
4. Note any changes or adjustments in the comments column

## Annex B - Acceptance Test
Annex B - Example Worksheet WCA Performance Acceptance Test (PAT)
Location |  |  | Station XYZ |  |  | Legend
Installation Type |  |  | Insertion |  |  | Input
Type of Sampling System |  |  | In Line |  |  | Calculation
Line Size (inches) |  |  | 12"
Service |  |  | LACT |  |  | Absolute Test Results | Total Water% | C. Sampler% | Sampler Test | WCA FWA % |  | WCA Test
WCA Technology |  |  | Microwave |  |  | Water Injection Test 1 | 0.4925 | 0.5036 | Pass | 0.504 |  | Pass
WCA ID |  |  | WCA 1 |  |  | Water Injection Test 2 | 0.4999 | 0.5036 | Pass | 0.504 |  | Pass
Sampling System ID |  |  | Metering System ABC Sampling
Product Group |  |  | WTI |  |  | Comments related to PAT here (if any)
Product used during PAT |  |  | ANS
Date of last PAT |  |  | 4519800%
Average Flow Rate (BPH) during PAT |  |  | 3000
Average Fluid Density (API) during PAT |  |  | 32.5
Complies with API MPMS 10.10 Section 6? |  |  | Yes
Lab Analytical Method used for PAT |  |  | Coulometric KF Mass
Baseline Start | [6 empty] | Baseline Start |  |  |  | Baseline Start
Sample # | Time | Composite Sample | WCA FWA | WCA Sample Avg | Manual Spot Avg |  | Manual Spot Sample Results |  |  |  | Composite Sample Results
| (hh:mm:ss) | Av Vol % Water | Vol % Water | Vol % Water | Vol % Water |  | Result #1 | Result #2 | Result #3 |  | Result #1 | 0.1
1 | 11:50 |  |  | 0.01 | 0.11567 |  | 0.11 | 0.125 | 0.112 |  | Result #2 | 0.11
2 | 12:00 |  |  | 0.15 | 0.11533 |  | 0.1192 | 0.1144 | 0.1124 |  | Result #3 | 0.1
3 | 12:10 |  |  | 0.11 | 0.10127 |  | 0.1037 | 0.0991 | 0.101
4 | 12:20 |  |  | 0.11 | 0.088733 |  | 0.0882 | 0.0893 | 0.0887
5 | 12:30 |  |  | 0.12 | 0.0946 |  | 0.0976 | 0.0926 | 0.0936
6 | 12:40 |  |  | 0.101 | 0.091933 |  | 0.0916 | 0.0906 | 0.0936
Averages |  | 0.10333 | 0.105 | 0.10017 | 0.10126
Water Injection 1 | [6 empty] | Water Injection 1 |  |  |  | Water Injection 1
Sample # | Time | Composite Sample | WCA FWA | WCA Sample Avg | Manual Spot Avg | Calc Water Injection | Manual Spot Sample Results |  |  |  | Composite Sample Results
| (hh:mm:ss) | Av Vol % Water | Vol % Water | Vol % Water | Vol % Water | Avg Vol % Water | Result #1 | Result #2 | Result #3 |  | Result #1 | 0.5
1 | 11:50 |  |  | 0.4999 | 0.49553 |  | 0.49 | 0.4978 | 0.4988 |  | Result #2 | 0.4999
2 | 12:00 |  |  | 0.4985 | 0.49573 |  | 0.4898 | 0.4985 | 0.4989 |  | Result #3 | 0.5109
3 | 12:10 |  |  | 0.5101 | 0.5109 |  | 0.51 | 0.5123 | 0.5104
4 | 12:20 |  |  | 0.5201 | 0.51967 |  | 0.52 | 0.5201 | 0.5189
5 | 12:30 |  |  | 0.5401 | 0.52863 |  | 0.53 | 0.5285 | 0.5274
6 | 12:40 |  |  | 0.5305 | 0.52033 |  | 0.52 | 0.5211 | 0.5199
Averages |  | 0.5036 | 0.504 | 0.51653 | 0.5118 | 0.55
Baseline Mid | [6 empty] | Baseline Mid |  |  |  | Baseline Mid
Sample # | Time | Composite Sample | WCA FWA | WCA Sample Avg | Manual Spot Avg |  | Manual Spot Sample Results |  |  |  | Composite Sample Results
| (hh:mm:ss) | Av Vol % Water | Vol % Water | Vol % Water | Vol % Water |  | Result #1 | Result #2 | Result #3 |  | Result #1 | 0.1
1 | 11:50 |  |  | 0.01 | 0.11567 |  | 0.11 | 0.125 | 0.112 |  | Result #2 | 0.11
2 | 12:00 |  |  | 0.15 | 0.11533 |  | 0.1192 | 0.1144 | 0.1124 |  | Result #3 | 0.1
3 | 12:10 |  |  | 0.11 | 0.10127 |  | 0.1037 | 0.0991 | 0.101
4 | 12:20 |  |  | 0.11 | 0.088733 |  | 0.0882 | 0.0893 | 0.0887
5 | 12:30 |  |  | 0.12 | 0.0946 |  | 0.0976 | 0.0926 | 0.0936
6 | 12:40 |  |  | 0.101 | 0.091933 |  | 0.0916 | 0.0906 | 0.0936
Averages |  | 0.10333 | 0.105 | 0.10017 | 0.10126
Water Injection 2 | [6 empty] | Water Injection 2 |  |  |  | Water Injection 2
Sample # | Time | Composite Sample | WCA FWA | WCA Sample Avg | Manual Spot Avg | Calc Water Injection | Manual Spot Sample Results |  |  |  | Composite Sample Results
| (hh:mm:ss) | Av Vol % Water | Vol % Water | Vol % Water | Vol % Water | Avg Vol % Water | Result #1 | Result #2 | Result #3 |  | Result #1 | 0.5
1 | 11:50 |  |  | 0.4999 | 0.49553 |  | 0.49 | 0.4978 | 0.4988 |  | Result #2 | 0.4999
2 | 12:00 |  |  | 0.4985 | 0.49573 |  | 0.4898 | 0.4985 | 0.4989 |  | Result #3 | 0.5109
3 | 12:10 |  |  | 0.5101 | 0.5109 |  | 0.51 | 0.5123 | 0.5104
4 | 12:20 |  |  | 0.5201 | 0.51967 |  | 0.52 | 0.5201 | 0.5189
5 | 12:30 |  |  | 0.5401 | 0.52863 |  | 0.53 | 0.5285 | 0.5274
6 | 12:40 |  |  | 0.5305 | 0.52033 |  | 0.52 | 0.5211 | 0.5199
Averages |  | 0.5036 | 0.504 | 0.51653 | 0.5118 | 0.55
Baseline End | [6 empty] | Baseline End |  |  |  | Baseline End
Sample # | Time | Composite Sample | WCA FWA | WCA Sample Avg | Manual Spot Avg |  | Manual Spot Sample Results |  |  |  | Composite Sample Results
| (hh:mm:ss) | Av Vol % Water | Vol % Water | Vol % Water | Vol % Water |  | Result #1 | Result #2 | Result #3 |  | Result #1 | 0.1
1 | 11:50 |  |  | 0.01 | 0.11567 |  | 0.11 | 0.125 | 0.112 |  | Result #2 | 0.11
2 | 12:00 |  |  | 0.15 | 0.11533 |  | 0.1192 | 0.1144 | 0.1124 |  | Result #3 | 0.1
3 | 12:10 |  |  | 0.11 | 0.10127 |  | 0.1037 | 0.0991 | 0.101
4 | 12:20 |  |  | 0.11 | 0.088733 |  | 0.0882 | 0.0893 | 0.0887
5 | 12:30 |  |  | 0.12 | 0.0946 |  | 0.0976 | 0.0926 | 0.0936
6 | 12:40 |  |  | 0.101 | 0.091933 |  | 0.0916 | 0.0906 | 0.0936
Averages |  | 0.10333 | 0.105 | 0.10017 | 0.10126

## Annex C - Ongoing Performance
Annex C - Water Cut Analyzer (WCA) Ongoing Verification Log for Product Groups (Maximum Permissible Error MPE) | [5 empty] | Batch Closing Date | Batch ID | Composite Sampler Result (%) | WCA FWA Result (%) | Error Measured - WCA vs. Sampling | Error Magnitude - WCA vs. Sampling | Error Magnitude - Running Mean | 2SD | Control Limit - Running (2SD) | Lower Control Limit - Running | MPE | WCA Pass / Fail? | Comments
[6 empty] | 4513900% | WTI - 2023-08-01 | 0.11% | 0.08% | -0.03% | 0.03% | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
Location |  | Station XYZ | No. of running batches (5 minimum) | 10 | Legend | 4514000% | WTI - 2023-08-02 | 0.23% | 0.27% | 0.04% | 0.04% | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
Installation Type |  | Insertion | Sampling and Analysis System MPE | 0.05% | Input | 4514100% | WTI - 2023-08-03 | 0.11% | 0.2% | 0.09% | 0.09% | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
Type of Sampling System |  | In Line | WCA MPE | 0.05% | Calculation | 4514200% | WTI - 2023-08-04 | 0.11% | 0.13% | 0.02% | 0.02% | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
Line Size (inches) |  | 12" | MPE (WCA vs. Sampling Uncertainty) | 0.07071% |  | 4514300% | WTI - 2023-08-05 | 0.11% | 0.09% | -0.02% | 0.02% | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
Service |  | LACT |  |  |  | 4514400% | WTI - 2023-08-06 | 0.11% | 0.15% | 0.04% | 0.04% | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
WCA Technology |  | Microwave |  |  |  | 4514500% | WTI - 2023-08-07 | 0.11% | 0.08% | -0.03% | 0.03% | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
WCA ID |  | WCA 1 |  |  |  | 4514600% | WTI - 2023-08-08 | 0.25% | 0.27% | 0.02% | 0.02% | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
Sampling System ID |  | Metering System ABC Sampling |  |  |  | 4514700% | WTI - 2023-08-09 | 0.11% | 0.15% | 0.04% | 0.04% | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
Product Group |  | WTI |  |  |  | 4514800% | WTI - 2023-08-10 | 0.11% | 0.13% | 0.02% | 0.02% | 0.035% | 0.04025% | 0.07525% | -0.005249% | 0.07071% | FAIL
[6 empty] | 4514900% | WTI - 2023-08-11 | 0.11% | 0.09% | -0.02% | 0.02% | 0.034% | 0.04118% | 0.07518% | -0.007183% | 0.07071% | FAIL | Batch finished earlier than expected
Notes: 1. Maximum permissible error (MPE's) for the sampling system and the water cut analyzer can be based on results of testing per API MPMS 8.2 or manufacturer's specified MPE for this application. 2. A water cut analyzer (WCA) is deemed suitable as a secondary measurement device while the primary sampling system is temporarily out of service, i.e., due to an equipment failure, when the magnitude of the measured error (WCA FWA vs. sampling results) is not greater than the maximum permissible error (MPE) for 10 consecutive batches. Thereafter to improve sensitivity the number of running points can be reduced to 5. The MPE is based on the uncertainty of the sampling and Analysis system (e.g., can be assumed as the max observed error during the water injection test under worst case conditions) and the MPE of the water cut analyzer (e.g., as advised by the manufacturer for application or otherwise statistically derived). The MPE is shown on the ongoing performance control chart below (red dashed line). 3. The ongoing performance control chart also includes control limits (green dashed line(s)). When a measured error is within the MPE and has exceeded these control limits, the cause should be investigated, including verification of the WCA. | [5 empty] | 4515000% | WTI - 2023-08-12 | 0.11% | 0.15% | 0.04% | 0.04% | 0.034% | 0.04118% | 0.07518% | -0.007183% | 0.07071% | FAIL
[6 empty] | 4515100% | WTI - 2023-08-13 | 0.11% | 0.12% | 0.01% | 0.01% | 0.026% | 0.0204% | 0.0464% | 0.005604% | 0.07071% | PASS
[6 empty] | 4515200% | WTI - 2023-08-14 | 0.11% | 0.13% | 0.02% | 0.02% | 0.026% | 0.0204% | 0.0464% | 0.005604% | 0.07071% | PASS
[6 empty] | 4515300% | WTI - 2023-08-15 | 0.11% | 0.09% | -0.02% | 0.02% | 0.026% | 0.0204% | 0.0464% | 0.005604% | 0.07071% | PASS
[6 empty] | 4515400% | WTI - 2023-08-16 | 0.11% | 0.15% | 0.04% | 0.04% | 0.026% | 0.0204% | 0.0464% | 0.005604% | 0.07071% | PASS
[6 empty] | 4515500% | WTI - 2023-08-17 | 0.11% | 0.08% | -0.03% | 0.03% | 0.026% | 0.0204% | 0.0464% | 0.005604% | 0.07071% | PASS
[6 empty] | 4515600% | WTI - 2023-08-18 | 0.3% | 0.27% | -0.03% | 0.03% | 0.027% | 0.0201% | 0.0471% | 0.0069% | 0.07071% | PASS
[6 empty] | 4515700% | WTI - 2023-08-19 | 0.11% | 0.15% | 0.04% | 0.04% | 0.027% | 0.0201% | 0.0471% | 0.0069% | 0.07071% | PASS
[6 empty] | 4515800% | WTI - 2023-08-20 | 0.11% | 0.13% | 0.02% | 0.02% | 0.027% | 0.0201% | 0.0471% | 0.0069% | 0.07071% | PASS
[6 empty] | 4515900% | WTI - 2023-08-21 | 0.11% | 0.09% | -0.02% | 0.02% | 0.027% | 0.0201% | 0.0471% | 0.0069% | 0.07071% | PASS
[6 empty] | 4516000% | WTI - 2023-08-22 | 0.11% | 0.15% | 0.04% | 0.04% | 0.027% | 0.0201% | 0.0471% | 0.0069% | 0.07071% | PASS
[6 empty] | 4516100% | WTI - 2023-08-23 | 0.11% | 0.14% | 0.03% | 0.03% | 0.029% | 0.01661% | 0.04561% | 0.01239% | 0.07071% | PASS
[6 empty] | 4516200% | WTI - 2023-08-24 | 0.11% | 0.15% | 0.04% | 0.04% | 0.031% | 0.01661% | 0.04761% | 0.01439% | 0.07071% | PASS
[6 empty] | 4516300% | WTI - 2023-08-25 | 0.11% | 0.08% | -0.03% | 0.03% | 0.032% | 0.01497% | 0.04697% | 0.01703% | 0.07071% | PASS
[6 empty] | 4516400% | WTI - 2023-08-26 | 0.11% | 0.15% | 0.04% | 0.04% | 0.032% | 0.01497% | 0.04697% | 0.01703% | 0.07071% | PASS
| Station XYZ - WTI - WCA 1 vs. Metering System ABC Sampling - Measured Error |  |  |  |  | 4516500% | WTI - 2023-08-27 | 0.11% | 0.16% | 0.05% | 0.05% | 0.034% | 0.01833% | 0.05233% | 0.01567% | 0.07071% | PASS
[6 empty] | 4516600% | WTI - 2023-08-28 | 0.11% | 0.13% | 0.02% | 0.02% | 0.033% | 0.0201% | 0.0531% | 0.0129% | 0.07071% | PASS
[6 empty] | 4516700% | WTI - 2023-08-29 | 0.11% | 0.09% | -0.02% | 0.02% | 0.031% | 0.02088% | 0.05188% | 0.01012% | 0.07071% | PASS
[6 empty] | 4516800% | WTI - 2023-08-30 | 0.11% | 0.15% | 0.04% | 0.04% | 0.033% | 0.0201% | 0.0531% | 0.0129% | 0.07071% | PASS
[6 empty] | 4516900% | WTI - 2023-08-31 | 0.11% | 0.08% | -0.03% | 0.03% | 0.034% | 0.01833% | 0.05233% | 0.01567% | 0.07071% | PASS
[6 empty] | 4517000% | WTI - 2023-09-01 | 0.11% | 0.16% | 0.05% | 0.05% | 0.035% | 0.02049% | 0.05549% | 0.01451% | 0.07071% | PASS
[6 empty] | 4517100% | WTI - 2023-09-02 | 0.11% | 0.15% | 0.04% | 0.04% | 0.036% | 0.0204% | 0.0564% | 0.0156% | 0.07071% | PASS
[6 empty] | 4517200% | WTI - 2023-09-03 | 0.11% | 0.13% | 0.02% | 0.02% | 0.034% | 0.02227% | 0.05627% | 0.01173% | 0.07071% | PASS
[6 empty] | 4517300% | WTI - 2023-09-04 | 0.11% | 0.09% | -0.02% | 0.02% | 0.033% | 0.02375% | 0.05675% | 0.009251% | 0.07071% | PASS
[6 empty] | 4517400% | WTI - 2023-09-05 | 0.11% | 0.15% | 0.04% | 0.04% | 0.033% | 0.02375% | 0.05675% | 0.009251% | 0.07071% | PASS
[6 empty] | 4517500% | WTI - 2023-09-06 | 0.11% | 0.08% | -0.03% | 0.03% | 0.031% | 0.02088% | 0.05188% | 0.01012% | 0.07071% | PASS
[6 empty] | 4517600% | WTI - 2023-09-07 | 0.11% | 0.07% | -0.04% | 0.04% | 0.033% | 0.0201% | 0.0531% | 0.0129% | 0.07071% | PASS
[6 empty] | 4517700% | WTI - 2023-09-08 | 0.11% | 0.15% | 0.04% | 0.04% | 0.035% | 0.01844% | 0.05344% | 0.01656% | 0.07071% | PASS
[6 empty] | 4517800% | WTI - 2023-09-09 | 0.11% | 0.13% | 0.02% | 0.02% | 0.033% | 0.0201% | 0.0531% | 0.0129% | 0.07071% | PASS
[6 empty] | 4517900% | WTI - 2023-09-10 | 0.11% | 0.09% | -0.02% | 0.02% | 0.032% | 0.02154% | 0.05354% | 0.01046% | 0.07071% | PASS
[6 empty] | 4518000% | WTI - 2023-09-11 | 0.11% | 0.15% | 0.04% | 0.04% | 0.031% | 0.01887% | 0.04987% | 0.01213% | 0.07071% | PASS
[6 empty] | 4518100% | WTI - 2023-09-12 | 0.11% | 0.08% | -0.03% | 0.03% | 0.03% | 0.01789% | 0.04789% | 0.01211% | 0.07071% | PASS
[6 empty] | 4518200% | WTI - 2023-09-13 | 0.11% | 0.15% | 0.04% | 0.04% | 0.032% | 0.01744% | 0.04944% | 0.01456% | 0.07071% | PASS
[6 empty] | 4518300% | WTI - 2023-09-14 | 0.11% | 0.14% | 0.03% | 0.03% | 0.033% | 0.01562% | 0.04862% | 0.01738% | 0.07071% | PASS
[6 empty] | 4518400% | WTI - 2023-09-15 | 0.11% | 0.13% | 0.02% | 0.02% | 0.031% | 0.01661% | 0.04761% | 0.01439% | 0.07071% | PASS
[6 empty] | 4518500% | WTI - 2023-09-16 | 0.11% | 0.09% | -0.02% | 0.02% | 0.03% | 0.01789% | 0.04789% | 0.01211% | 0.07071% | PASS
[6 empty] | 4518600% | WTI - 2023-09-17 | 0.11% | 0.15% | 0.04% | 0.04% | 0.03% | 0.01789% | 0.04789% | 0.01211% | 0.07071% | PASS
| Station XYZ - WTI - WCA 1 vs. Metering System ABC Sampling - Ongoing Performance Control Chart |  |  |  |  | 4518700% | WTI - 2023-09-18 | 0.11% | 0.08% | -0.03% | 0.03% | 0.029% | 0.01661% | 0.04561% | 0.01239% | 0.07071% | PASS
[6 empty] | 4518800% | WTI - 2023-09-19 | 0.11% | 0.16% | 0.05% | 0.05% | 0.032% | 0.0196% | 0.0516% | 0.0124% | 0.07071% | PASS
[6 empty] | 4518900% | WTI - 2023-09-20 | 0.11% | 0.15% | 0.04% | 0.04% | 0.034% | 0.01833% | 0.05233% | 0.01567% | 0.07071% | PASS
[6 empty] | 4519000% | WTI - 2023-09-21 | 0.11% | 0.13% | 0.02% | 0.02% | 0.032% | 0.0196% | 0.0516% | 0.0124% | 0.07071% | PASS
[6 empty] | 4519100% | WTI - 2023-09-22 | 0.11% | 0.09% | -0.02% | 0.02% | 0.031% | 0.02088% | 0.05188% | 0.01012% | 0.07071% | PASS
[6 empty] | 4519200% | WTI - 2023-09-23 | 0.11% | 0.15% | 0.04% | 0.04% | 0.031% | 0.02088% | 0.05188% | 0.01012% | 0.07071% | PASS
[6 empty] | 4519300% | WTI - 2023-09-24 | 0.11% | 0.08% | -0.03% | 0.03% | 0.031% | 0.02088% | 0.05188% | 0.01012% | 0.07071% | PASS
[6 empty] | 4519400% | WTI - 2023-09-25 | 0.11% | 0.16% | 0.05% | 0.05% | 0.034% | 0.02227% | 0.05627% | 0.01173% | 0.07071% | PASS
[6 empty] | 4519500% | WTI - 2023-09-26 | 0.11% | 0.15% | 0.04% | 0.04% | 0.036% | 0.0204% | 0.0564% | 0.0156% | 0.07071% | PASS
[6 empty] | 4519600% | WTI - 2023-09-27 | 0.11% | 0.13% | 0.02% | 0.02% | 0.034% | 0.02227% | 0.05627% | 0.01173% | 0.07071% | PASS
[6 empty] | 4519700% | WTI - 2023-09-28 | 0.11% | 0.09% | -0.02% | 0.02% | 0.033% | 0.02375% | 0.05675% | 0.009251% | 0.07071% | PASS
[6 empty] | 4519800% | WTI - 2023-09-29 | 0.11% | 0.15% | 0.04% | 0.04% | 0.032% | 0.02154% | 0.05354% | 0.01046% | 0.07071% | PASS
[6 empty] | 4519900% | WTI - 2023-09-30 | 0.11% | 0.08% | -0.03% | 0.03% | 0.031% | 0.02088% | 0.05188% | 0.01012% | 0.07071% | PASS
[6 empty] | 4520000% | WTI - 2023-10-01 | 0.11% | 0.15% | 0.04% | 0.04% | 0.033% | 0.0201% | 0.0531% | 0.0129% | 0.07071% | PASS
[6 empty] | 4520100% | WTI - 2023-10-02 | 0.11% | 0.16% | 0.05% | 0.05% | 0.036% | 0.0204% | 0.0564% | 0.0156% | 0.07071% | PASS
[6 empty] | 4520200% | WTI - 2023-10-03 | 0.11% | 0.13% | 0.02% | 0.02% | 0.034% | 0.02227% | 0.05627% | 0.01173% | 0.07071% | PASS
[6 empty] | 4520300% | WTI - 2023-10-04 | 0.11% | 0.09% | -0.02% | 0.02% | 0.033% | 0.02375% | 0.05675% | 0.009251% | 0.07071% | PASS
[6 empty] | 4520400% | WTI - 2023-10-05 | 0.11% | 0.15% | 0.04% | 0.04% | 0.032% | 0.02154% | 0.05354% | 0.01046% | 0.07071% | PASS
[6 empty] | 4520500% | WTI - 2023-10-06 | 0.11% | 0.08% | -0.03% | 0.03% | 0.031% | 0.02088% | 0.05188% | 0.01012% | 0.07071% | PASS
[6 empty] | 4520600% | WTI - 2023-10-07 | 0.11% | 0.13% | 0.02% | 0.02% | 0.031% | 0.02088% | 0.05188% | 0.01012% | 0.07071% | PASS
[6 empty] | 4520700% | WTI - 2023-10-08 | 0.11% | 0.15% | 0.04% | 0.04% | 0.033% | 0.0201% | 0.0531% | 0.0129% | 0.07071% | PASS
[6 empty] | 4520800% | WTI - 2023-10-09 | 0.11% | 0.13% | 0.02% | 0.02% | 0.031% | 0.02088% | 0.05188% | 0.01012% | 0.07071% | PASS
[6 empty] | 4520900% | WTI - 2023-10-10 | 0.11% | 0.09% | -0.02% | 0.02% | 0.03% | 0.02191% | 0.05191% | 0.008091% | 0.07071% | PASS
[6 empty] | 4521000% | WTI - 2023-10-11 | 0.11% | 0.15% | 0.04% | 0.04% | 0.03% | 0.02191% | 0.05191% | 0.008091% | 0.07071% | PASS
[6 empty] | 4521100% | WTI - 2023-10-12 | 0.11% | 0.13% | 0.02% | 0.02% | 0.027% | 0.018% | 0.045% | 0.009% | 0.07071% | PASS
[6 empty] | 4521200% | WTI - 2023-10-13 | 0.11% | 0.16% | 0.05% | 0.05% | 0.03% | 0.02191% | 0.05191% | 0.008091% | 0.07071% | PASS
[6 empty] | 4523000% | WTI - 2023-10-31 | 0.11% | 0.17% | 0.06% | 0.06% | 0.034% | 0.02713% | 0.06113% | 0.006871% | 0.07071% | PASS
Note on AC1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Pass if previous 10 consecutive results are within MPE.